## Preliminary results
[5 empty] | fluka |  | calc
| Beam Energy | LET [MeV cm2/g) | Ratio | MeV/mm | Ca Dose/p [Gy] | Ca Dose/p [rad] | Dose /p rad | CS |  | No Bits/FF/PLL | Rate Max [SEU/s/bit or PLL] | MTBF [s] | Rate Max [SEU/s/bit or PLL] | MTBF [s] | Max dose [krad/min] | Max dose rate [krad/h] | Min dose [krad/min] | Min dose rate [krad/h]
PLL-OCL | 22 | 19.56 | 1 | 4.5575 | 2.94e-10 | 2.94e-08 | 3.1335e-07 | 9.5e-11 | /pll | 1 | 0.02375 | 42.1053 | 0.000152 | 6578.9474 | 4.7003 | 282.0161 | 0.030082 | 1.8049
PLL-Tullio Grassi | 230 | 3.432 | 0.17546 | 0.79966 | 5.1585e-11 | 5.1585e-09 | 5.4981e-08 | 2e-10 | /pll | 1 | 0.05 | 20 | 0.00032 | 3125 | 0.82471 | 49.4826 | 0.0052781 | 0.31669
FF-Praha | 37 | 12.95 | 0.66207 | 3.0173 | 1.9465e-10 | 1.9465e-08 | 2.0746e-07 | 1.5e-14 | /bit | 2000 | 3.75e-06 | 133.3333 | 2.4e-08 | 20833.3333 | 3.1119 | 186.7131 | 0.019916 | 1.195
FF-Praha | 28 | 16.06 | 0.82106 | 3.742 | 2.4139e-10 | 2.4139e-08 | 2.5728e-07 | 1.22e-14 | /bit | 2000 | 3.05e-06 | 163.9344 | 1.952e-08 | 25614.7541 | 3.8592 | 231.5531 | 0.024699 | 1.4819
uSRAM-OCL | 22 | 19.56 | 1 | 4.5575 | 2.94e-10 | 2.94e-08 | 3.1335e-07 | 1.1e-14 | /bit | 65536 | 2.75e-06 | 5.5487 | 1.76e-08 | 866.9767 | 4.7003 | 282.0161 | 0.030082 | 1.8049
lSRAM-OCL | 22 | 19.56 | 1 | 4.5575 | 2.94e-10 | 2.94e-08 | 3.1335e-07 | 2e-14 | /bit | 1114112 | 5e-06 | 0.17952 | 3.2e-08 | 28.0492 | 4.7003 | 282.0161 | 0.030082 | 1.8049
TSL | 180 | 4.007 | 0.20486 | 0.93363 | 6.0228e-11 | 6.0228e-09 | 6.4192e-08 | [7 empty] | 0.96288 | 57.7729 | 0.0061624 | 0.36975
[14 empty] | Max dose in 24 hrs | 1386.5502
Flux Max | 250000000
Flux min | 1600000
| #SEU | CS/bit or PLL | Bit | Flux | Time [s] | Time [h] |  | Dose kRad
200
PLL | 200 | 2e-10 | 1 | 250000000 | 4000 | 1.1111 |  | 6.0228 |  | 230 MeV
FF | 200 | 1.5e-14 | 2000 | 250000000 | 26666.6667 | 7.4074 |  | 40.1519 |  | 37 MeV
uSRAM | 1000 | 1.1e-14 | 65536 | 250000000 | 5548.6506 | 1.5413 |  | 8.3546 |  | 22 MeV
lSRAM | 1000 | 2e-14 | 1114112 | 250000000 | 179.5152 | 0.049865 |  | 0.2703 |  | 22 MeV
Assuming some increase in CS due to energy will also decrease the total dose.
FF Design |  | 40 | 80 | 160
inv/Freq
| 0 | 100 | 100 | 100
| 4 | 100 | 100 | 100
| 8 | 100 | 100 | 100
Run status
Calibration |  | 78500 |  |  |  | 1e-10
cnt 10 |  | 35000
Flux |  | 1990942.029
fluence |  | 2747500000
sec |  | 1380
hours |  | 0.38333
Dose [rad] |  | 16.5476
Flux |  | 100000000
CS |  | 1e-10
Target error |  | 100
Time to target |  | 10000 | s
|  | 2.7778 | hr
Target dose |  | 10000 | rad
Dose/s |  | 0.60228 | rad
Time to target |  | 16603.5968 | s
|  | 4.6121 | hr
|  |  | 1.602e-07 | 0.001833 | 2.9365e-10

## Sheet1
[7 empty] | sum | [5 empty] | SEL
Date | start | stop | cnt10 | cnt12 (100Hz) | Time [s] | Beam Hz | Calibration | Fluence | Flux | Dose rad/p | Dose krad | log file |  |  | RCU2 comment | Comment
20140507 | 2:40 | 3:05 | 35431 | 134000 | 1340 | 1 | 78500 | 2781333500 | 2075622.0149 | 6e-08 | 0.16688 | 20140507-0245-rcu2-serdes-r1.txt |  |  |  | Beam not on for the first 220 sec
20140507 | 3:25 | 3:32 | 177453 | 14235 | 142.35 | 50 | 78500 | 13930060500 | 97857818.7566 | 6e-08 | 0.8358 | 20140507-0325-rcu2-serdes-r2.txt |  |  |  | Increase beam by x50
20140507 | 3:44 | 3:48 | 252633 | 16000 | 160 | 50 | 78500 | 19831690500 | 123948065.625 | 6e-08 | 1.1899 | 20140507-0344-rcu2-serdes-r3.txt |  |  | 03:46 link down
20140507 | 4:00 | 4:14 | 1356473 | 83854 | 838.54 | 50 | 78500 | 106483130500 | 126986345.9107 | 6e-08 | 6.389 | 20140507-0400-rcu2-serdes-r4.txt |  |  | 0407 bit error,0412 bit error,0414 link down stop run | Beam after roughly 10 seconds after pressing bean enable
20140507 | 4:20 | 4:23 | 338464 | 21260 | 212.6 | 50 | 78500 | 26569424000 | 124973772.3424 | 6e-08 | 1.5942 | 20140507-0420-rcu2-serdes-r5.txt |  |  |  | Beam on after 15 s, turned of to increase resolution of Voltage
20140507 | 4:26 | 4:35 | 857977 | 53453 | 534.53 | 50 | 78500 | 67351194500 | 126000775.4476 | 6e-08 | 4.0411 | 20140507-0426-rcu2-serdes-r6.txt |  |  |  | Beam on after 9 sec,turned off to increase threshold again
20140507 | 4:38 | 4:48 | 1028015 | 63151 | 631.51 | 50 | 78500 | 80699177500 | 127787647.8599 | 6e-08 | 4.842 | 20140507-0438-rcu2-serdes-r7.txt |  |  | 4:38 bit error,4:45 high jump and reset of counter, lost connection to computer, stopped beam | Beam on after 16 s, after this run the 1V2 is at 1.5mV after repower of rcu2 but down again to 1.45mV after some seconds
20140507 | 5:10 | 5:55 | 4587151 | 277505 | 2775.05 | 50 | 78500 | 360091353500 | 129760311.8863 | 6e-08 | 21.6055 | 20140507-0510-rcu2-optical-r8.txt |  |  | 05:12 bit error with loss of link afterwards, repowered and link is up. 05:!2 link down repowerd link up, 05:15 bit error, 05:16 link down and up again automatically followed ,by bit error and a reset of counter, 05:19 link down -> repower -> reset counter, 05:22 link down-> repower->link up,05:20 no errors for a long while -> repower board, 05:43 bit error noticed, 05:56 stopped beam (noticed that pc in hall runs on battery | Moved beam to optical link, beam on after 8 s
20140508 | 9:52 | 9:54 | 14641 | 5207 | 52.07 | 11 | 78500 | 1149318500 | 22072565.7768 | 6e-08 | 0.068959 | 20140508-0952-rcu2-trigger-r1.txt |  |  | 1s trigger rate | First run with trigger, beam on after 10 s, beam stopped due to problems on line
20140508 | 9:56 | 10:02 | 19768 | 7364 | 73.64 | 11 | 78500 | 1551788000 | 21072623.5741 | 6e-08 | 0.093107 | 20140508-0956-rcu2-trigger-r2.txt |  |  | 1s trigger rate | Beam on after 15s, we stopped the beam as the gui did no longer read values
20140508 | 10:09 | 10:40 | 621703 | 183654 | 1836.54 | 11 | 78500 | 48803685500 | 26573712.2524 | 6e-08 | 2.9282 | 20140508-1009-rcu2-trigger-r3.txt | 42 | 8.3407e-10 | 1s trigger rate | Repowerd after >220s as no contact with, repowered everytime we lost connection. Keep one continous log file. 10:40 stopped run when shifting to beam therapy
20140508 |  |  | 62101 |  |  | 11 | 78500 | 4874928500 | #DIV/0! | 6e-08 | 0.2925 |  |  |  |  | Some fluence received in the process of switching to therapy. Problem with the beam controlling program.
20140508 | 11:34 | 11:38 | 44720 | 15174 | 151.74 | 11 | 78500 | 3510520000 | 23135099.5123 | 6e-08 | 0.21063 | 20140508-1134-rcu2-trigger-r4.txt |  |  | 1s trigger rate | Repowered in the beginning(probably forgot from end of previous run, 10:38 switch to teraphy keep running,
20140508 | 11:48 | 12:19 | 636561 | 189325 | 1893.25 | 11 | 78500 | 49970038500 | 26393787.6667 | 6e-08 | 2.9982 | 20140508-1134-rcu2-trigger-r4.txt | 28 | 5.2355e-10 | 1s trigger rate, 11:50 high increase in pll counting->repowerd to fix this problem, | We reset the beam counters during proton therapy so added a new line (same log files),12:20 stopped to increase trigger rate
20140508 | 12:20 | 12:59 | 748575 | 220737 | 2207.37 | 11 | 78500 | 58763137500 | 26621335.5713 | 6e-08 | 3.5258 | 20140508-1221-rcu2-trigger-r5.txt | 38 | 6.4666e-10 | 1ms trigger rate,12:25 possible bananas tigger counters -> repower | stopped run at 12:59 to change trigger rate (also take lunch as proton therapy starts)
20140508 | 13:45 | 14:02 | 300903 | 96000 | 960 | 11 | 78500 | 23620885500 | 24605089.0625 | 6e-08 | 1.4173 | 20140508-1345-rcu2-trigger-r6.txt | 20 | 8.4671e-10 | 0.1ms trigger | beam start after 30 s, stopped beam to reposition to optical link
Probably run for ~5min with beam but without logging due to software problem when switching to therapy. Estimated additional dose ~0,5kRads
20140508 | 14:45 | 15:13 | 445841 | 186472 | 1864.72 | 11 | 78500 | 34998518500 | 18768779.495 | 6e-08 | 2.0999 | 20140508-1445-rcu2-ttc-r1.txt |  |  | 1s trigger rate | Beam available after ~50 s in current log file., 14:47 shift beam to p therapy,15:05 beam back in blue hall ,
20140508 | 15:15 | 15:32 | 348743 | 104321 | 1043.21 | 11 | 78500 | 27376325500 | 26242391.7524 | 6e-08 | 1.6426 | 20140508-1515-rcu2-ttc-r2.txt |  |  | 1ms trigger rate, new trigger rate to see if hamming errors increase.
20140508 | 15:33 | 15:57 | 479039 | 142258 | 1422.58 | 11 | 78500 | 37604561500 | 26434057.4871 | 6e-08 | 2.2563 | 20140508-1533-rcu2-ttc-r3.txt |  |  | 200hz tigger rate | 15:50 repowered due to decrease updating frequency in the software gui-> back to normal frequency after repowering.
Starterkit 1
20140508 | 1915 | 19:20 | 101421 | 30866 | 308.66 | 11 | 78500 | 7961548500 | 25793910.7756 | 6e-08 | 0.47769 | 20140-1912-sk-r1.txt | 4 | 2.0097e-13 | l2500 inv4 f40 | beam on after 17 s,stopped to fix resetting of counters.
20140508 | 19:24 | 19:54 | 594231 | 176528 | 1765.28 | 11 | 78500 | 46647133500 | 26424778.7886 | 6e-08 | 2.7988 | 20140-1924-sk-r2.txt | 3 | 2.5725e-14 |  | 19:29 reset after lots of errors in sram,19:32 current decrease->labview no function-> repower,19:54 stopped to increase beam intensity (also added delay after repowering to see if this helps for the resetting of the plls)
20140508 | 19:56 | 20:16 | 1046147 | 85671 | 856.71 | 11 | 78500 | 82122539500 | 95858037.7257 | 6e-08 | 4.9274 | 20140-1956-sk-r3.txt | 5 | 2.4354e-14 |  | 20:17 stopped to increase frequency
20140508 | 20:18 | 20:38 | 1201327 | 99494 | 994.94 | 11 | 78500 | 94304169500 | 94783775.4035 | 6e-08 | 5.6583 | 20140-2020-sk-r4.txt | 5 | 2.1208e-14 | l2500 inv4 f80 | 20:38 stopped beam due to problems with current
Starterkit 2 | Change resistors to 0x025
20140508 | 22:15 | 22:18 | 78895 | 8428 | 84.28 |  | 78500 | 6193257500 | 73484308.2582 | 6e-08 | 0.3716 | 20140-2215-sk2-r1.txt | 2 | 1.2917e-13 | L2500 inv0 f160 | Beam stopped due to problemes with current -> needed to go into the hall and repower.
20140508 | 22:22 | 22:25 | 102834 | 9290 | 92.9 |  | 78500 | 8072469000 | 86894176.5339 | 6e-08 | 0.48435 | 20140-2223-sk2-r2.txt | 1 | 4.9551e-14 |  | Stopped again due to problems with resetting power -> acces to see if reg enable is proberly connected., forgot to solder pin
20140508 | 22:48 | 22:59 | 671613 | 55897 | 558.97 |  | 78500 | 52721620500 | 94319230.9068 | 6e-08 | 3.1633 | 20140-2250-sk2-r3.txt | 3 | 2.2761e-14 |  | Stopped beam due to problems with current, access to repower
20140508 | 23:07 | 23:48 | 2881340 | 238317 | 2383.17 |  | 78500 | 226185190000 | 94909381.2023 | 6e-08 | 13.5711 | 20140-2308-sk2-r4.txt | 10 | 1.7685e-14 | run different frequencies | 23:25 Changed to 80MHz after 943 s 23:37 changed to 40Mhz after 1671s,23:48 stopped beam for the night
RCU Trigger with lower resistor value
20140509 | 9:39 | 10:24 | 1602223 | 128521 | 1285.21 | 11 | 78500 | 125774505500 | 97862999.432 | 6e-08 | 7.5465 | 20140509-0950-trigger-r7.txt | 6 | 4.7704e-11 | 1khz, triggers are not working | Beam on after 25 s in log file, stopped beam for proton therapy
20140509 | 10:35 | 10:49 | 409828 | 87737 | 877.37 | 11 | 78500 | 32171498000 | 36668108.0958 | 6e-08 | 1.9303 | 20140509-1035-trigger-r8.txt | 5 | 1.5542e-10 | 1khz, triggers are not working
20140509 | 11:00 | 11:31 | 3501244 | 188734 | 1887.34 | 11 | 78500 | 274847654000 | 145626995.6659 | 6e-08 | 16.4909 | 20140509-1100-trigger-r9.txt | 23 | 8.3683e-11 | 11:02 current decreased to 1.7mV do not know why, 11:18 repowering seems to bring the current down to 2.9mV only and not 2mV. 11:22 problems repowering could be software catching of character,
ESA | [17 empty] | CS | Bits
[13 empty] | DIE0 | Die1 | Die2 | Die3 | Total
20140509 | 12:30 | 12:34 | 54772 | 18104 | 181.04 | 11 | 78500 | 4299602000 | 23749458.6832 | 6e-08 | 0.25798 | 20140409-tsl-r1.csv | 460 | 427 | 440 | 454 | 1781 | 2.469e-14 | 16777216 | 10 Hz
20140509 | 12:39 | 12:41 | 145594 | 12059 | 120.59 | 11 | 78500 | 11429129000 | 94776755.9499 | 6e-08 | 0.68575 | 20140409-tsl-r2.csv | 1050 | 1120 | 1131 | 1137 | 4438 | 2.3145e-14 | 16777216 | 40 Hz
20140509 | 12:43 | 12:44 | 73707 | 6091 | 60.91 | 11 | 78500 | 5785999500 | 94992603.8417 | 6e-08 | 0.34716 | 20140409-tsl-r3.csv | 575 | 528 | 600 | 547 | 2250 | 2.3178e-14 | 16777216 | 40 Hz
20140509 | 12:46 | 12:49 | 45982 | 15032 | 150.32 | 11 | 78500 | 3609587000 | 24012686.2693 | 6e-08 | 0.21658 | 20140409-tsl-r4.csv | 327 | 365 | 334 | 346 | 1372 | 2.2656e-14 | 16777216 | 20 Hz
Sk ES dose test
20140509 | 14:46 |  | 236159 | 20007 | 200.07 | 11 | 78500 | 18538481500 | 92659976.5082 | 6e-08 | 1.1123 | 20140509-skdosetest-r1.txt |  |  | Reprogramming at start Ok
|  |  | 501203 |  | 0 | 11 | 78500 | 39344435500 | #DIV/0! | 6e-08 | 2.3607 |  |  |  | 15:06: switchig to therapy,programming OK
|  |  | 749227 |  |  | 11 | 78500 | 58814319500 | #DIV/0! | 6e-08 | 3.5289 |  |  |  | One failed programming, two failed programming, DONE!
| 2 | [6 empty] | 2e-14 | 10000 | 100000000 | 0.02 | 50
5
5
|  |  |  | cycles | length | frequency | period | Time readout
|  |  |  | 19453 | 128 | 20000000 | 5e-08 | 0.12 | 0.1245

## Sheet2
| Conv | 1 | eV | 1.602e-19 | J
LET
Dose | dE/dx * flux | = | (MeVcm2/mg)*cm2
Dose(krad) | 1,6e-19*1e6*100/(1e-6*1000) | 1.6e-08 | *LET